## Slide 1: Collective review Cardiac output monitoring
R3 ธีรวัฒน์ คุณาธินันท์
อาจารย์ ณรงค์ศักดิ์ เจษฎาภัทรกุล

## Slide 2: Why monitor cardiac output ?

## Slide 5: Fick's Principle
In 1870, Dr. Adolf Fick discovered a method of computing an animal’s CO based on the oxygen consumption and the difference in oxygen content between arterial and venous blood

## Slide 6: Fick's Principle
The principle:
" the total uptake of (or release of) a substance by the peripheral tissues is equal to the product of the blood flow to the peripheral tissues and the arterial-venous concentration difference (gradient) of the substance."

## Slide 8: Cardiac output monitoring
Invasive cardiac output monitoring
Semi/minimally invasive cardiac output monitoring
Non-invasive cardiac output monitoring

## Slide 9: Cardiac output monitoring
Invasive | Semi- invasive | Non -invasive
Continuous cardiac output monitoring (CCO) via PAC | FloTrac | Volume clamp methods
Intermittent thermodilution via PAC | Volume View | Bioreactance (NICOM)
 | PRAM | Bioimpedance
 | PiCCO | 
 | LidCO | 
 | Partial CO2 rebreathing technique | 
 | Transesophageal echocardiogram | 
 | Esophageal doppler monitor | 

## Slide 10: Invasive cardiac output monitoring

## Slide 11: Pulmonary artery catheter

## Slide 12: Pulmonary artery catheter

## Slide 13: Pulmonary artery catheter

## Slide 14: Pulmonary artery catheter
Advantages
  The standard method for CO measurement
  PAC provides, in addition, PA pressures, PAOP,mixed-venous oxygen saturation, RVEF and RVEDV

## Slide 15: Pulmonary artery catheter
Limitations
  Complications associated with placement and presence of a PAC

## Slide 18: Semi-invasive cardiac output monitoring

## Slide 19: Pulse contour analysis
Stroke volume (SV)and systemic vascular resistance (SVR) calculate from arterial waveform

## Slide 20: PiCCO
PiCCO monitor (Pulse Medical System, Munich, Germany)
PiCCO applies a special algorithm that combines real-time continuous monitoring through pulse contour analysis with intermittent transpulmonary thermodilution (TPTD) measurement

## Slide 23: PiCCO

## Slide 25: PiCCO
PiCCO arterial contour analysis uses the TPTD technique as an external calibration
The calibration interval is recommended to be every 8 h or whenever there is a clinically significant change in SVR

## Slide 26: PiCCO
Advantages
  less invasive
  useful in the pediatric population when a PAC is too large to be inserted
  independent of ventilator and respiratory cycles
  measure global end-diastolic volume (GEDV) and intrathoracic blood volume (ITBV)

## Slide 27: PiCCO

## Slide 28: PiCCO

## Slide 29: PiCCO

## Slide 33: PiCCO
Limitations
  regurgitant valve
  intracardiac shunt
  extracorporeal circulation

## Slide 34: LiDCO
LiDCOplus system (LiDCO, Cambridge, UK)
based on running two proprietary algorithms
  indicator dilution CO monitoring (LiDCO system)
  continuous arterial waveform analysis (PulseCO system)
Consists of a lithium sensor attached to the arterial line
LiDCO system requires only an arterial line and a peripheral IV line

## Slide 36: LiDCO
lithium is injected into the venous circulation, blood samples from the arterial line are drawn, and a lithium concentration time curve is plotted
area under the curve will determine CO
lithium indicator can be injected via either central or peripheral venous access

## Slide 37: LiDCO
LiDCOplus system provides various parameters
  CO
  intrathoracic blood volume (ITBV)
  MAP, SVR, SV
  SVV and pulse pressure variation (PPV)
PulseCO system is recommended to be recalibrated every 8 h or with each major hemodynamic change

## Slide 38: LiDCO

## Slide 39: LiDCO
Advantages
  less invasive than PAC and PiCCO
Limitations
  aortic regurgitation
  severe arrhythmia
  severe peripheral vasoconstriction
  patients who receive lithium therapy

## Slide 40: LiDCO
Contraindications
  conditions related to a patient’s extra lithium intake because this will lead to an overestimate of CO
  patients who receive nondepolarizing muscle relaxant which will interfere with the lithium sensor
  anatomic cardiac abnormalities
    aortic valve regurgitation
    intraaortic balloon pump (IABP)
    poor quality of arterial signal

## Slide 43: Volume View
The VolumeView system determines CO by continuous arterial pressure analysis on the femoral artery and external calibration using the TPTD technique
system consists of
  specific thermistortipped arterial catheter (the VolumeView catheter) and the EV1000 monitoring platform
  continuous central venous oxygen saturation (ScvO2) monitoring via the PreSep oximetry catheter

## Slide 44: Volume View

## Slide 50: Volume view
Provide Informative volumetric parameters
  Extravascular Lung Water (EVLW)
  Pulmonary Vascular Permeability Index (PVPI)
  Global End-Diastolic Volume (GEDV)
  Global Ejection Fraction (GEF)
  Cardiac function index (CFI)

## Slide 54: GEDV- Global End Diastolic Volume

## Slide 55: GEDV- Global End Diastolic Volume

## Slide 56: GEDV- Global End Diastolic Volume

## Slide 57: GEDV- Global End Diastolic Volume

## Slide 58: EVLW-Extra Vascular Lung Water

## Slide 59: EVLW-Extra Vascular Lung Water

## Slide 60: EVLW-Extra Vascular Lung Water

## Slide 61: EVLW-Extra Vascular Lung Water

## Slide 62: Limitations
Improper transducer position
Over/under-damped pressure lines
Arterial pressure inaccurate/not representative of aortic pressure
Excessive patient movement
Incorrect placement or position of the femoral artery catheter
Intra-aortic balloon pumps
Anatomical abnormalities (for example, cardiac shunts)

## Slide 63: Limitations
Excessive variations/interference in blood temperature measurement such as
  Post cardiopulmonary bypass surgery status
  Centrally administered cooled/warmed solutions of blood products
  Clot formation on the thermistor
  External heat sources
    Electrocautery or electrosurgical unit interference
    Rapid changes in cardiac output

## Slide 64: Applications in ICU
Determine possible causes of hemodynamic instability
  Severe sepsis/Septic shock
  ARDS/Acute lung injury
  Pulmonary edema
  Cardiogenic shock
  Severe burn

## Slide 65: Applications in ICU
Provide valuable hemodynamic information
  Liver surgery
  Cardiac surgery
  Transplantation

## Slide 66: FloTrac
Derive SV from arterial waveform, after correcting for compliance and resistance of arterial system
Continuous computes stroke volume from arterial pressure signal

## Slide 70: FloTrac
Connect to existing arterial catheter
Require no additional catheters, boluses or chemical indicators
Continuously monitors changes in the patient’s vascular tone (compliance and resistance)
Compatible with Vigileo monitor and the EV1000 clinical platform

## Slide 71: FloTrac
Advantages
  less invasive
  provides continuous
  CO monitoring
  Relatively easy to use

## Slide 72: FloTrac
Limitations
  Improper transducer posiotion
  Over/under-damped pressure lines
  Some conditions causing BP variations
    Intra-aortic balloon pumps
  Excessive patient movement
  Electrocautery or electrosurgical unit interference

## Slide 73: FloTrac
Limitations
  Arterial pressure deemed inaccurate or not representative of aortic pressure, including
    Extreme peripheral vasoconstriction
      Compromised radial arterial pressure waveform
    Hyperdynamic conditions
      seen in post liver transplant
  Aortic valve regurgitation - cause an over estimation of Stroke Volume / Cardiac Output

## Slide 74: PRAM
PRAM (pressure‑recording analytical method)
  arterial pressure-derived continuous CO measurement
  PRAM needs only an arterial line as with FloTrac/Vigileo
Hemodynamic parameters including CO, SVV, PPV, and SVR

## Slide 75: PRAM
principle that, in any given vessel, volume changes occur mainly because of radial expansion in response to pressure variations; simply put, the alterations of the systolic portion of the area under the curve reflect changes in SV

## Slide 76: PRAM
Advantages
  less invasive technique
  continuous monitoring of CO
  advanced hemodynamic parameters including SVV and PPV

## Slide 77: PRAM
Limitations
  technically related (over-damping or under-damping of arterial waveforms)
  inappropriate signal acquisition (e.g., aortic valve regurgitation)
  abnormality of the peripheral arteries (e.g., aortic dissection, atherosclerotic plaque)

## Slide 78: Transesophageal echocardiogram (TEE)

## Slide 79: Indications
Intraoperative TEE
  All open heart (i.e., valvular) and thoracic aortic surgical procedures
  Use in some coronary artery bypass graft surgeries
  Noncardiac surgery when patients have known or suspected cardiovascular pathology which may impact outcomes

## Slide 80: Indications
Critically ill patients
  Patients in whom diagnostic information is not obtainable by TTE and this information is expected to alter management

## Slide 81: Indications
Evaluation of cardiac and aortic structure and function in situations where the findings will alter management
Guidance of transcatheter procedures

## Slide 82: Transesophageal echocardiogram (TEE)
Advantages
  detect anatomical abnormalities
  volume status
  myocardial contractility information and other functional assessment as well as hemodynamic parameters
  real-time measurement of CO

## Slide 83: Transesophageal echocardiogram (TEE)
Limitations
  anesthetized patients
  cannot be used in very small children
  accuracy is also highly dependent upon the quality of echocardiographic images and the operator’s skill and experience

## Slide 84: Absolute contraindication
Perforated viscus
Esophageal stricture
Esophageal tumor
Esophageal perforation, laceration
Esophageal diverticulum
Active upper GI bleed

## Slide 85: Relative contraindication
History of radiation to neck and mediastinum
History of GI surgery
Recent upper GI bleed
Barrett’s esophagus
History of dysphagia
Symptomatic hiatal hernia

## Slide 86: Relative contraindication
Restriction of neck mobility
  Severe cervical arthritis
  Atlantoaxial joint disease
Esophageal varices
Coagulopathy, thrombocytopenia
Active esophagitis
Active peptic ulcer disease

## Slide 87: Esophageal doppler monitor
Doppler effect
  Measure velocity of blood in Descending aorta in cm/s
  Diameter of the aorta
  Base on age, sex, weight,height

## Slide 88: Esophageal doppler monitor

## Slide 92: Esophageal doppler monitor
Parameter
  CO ,CI
  SV ,SVV
  FTc

## Slide 93: Esophageal doppler monitor

## Slide 94: Esophageal doppler monitor

## Slide 95: Limitations
Aortic pathology or compression or due to abnormal upper to lower body blood flow distribution
Aortic cross-sectional area - Not constant
  Change in pulse pressure, vascular tone, aortic compliance, volume status, or catecholamine use

## Slide 96: Limitations
Probe position - Critical to obtain accurate measurement for blood flow and aortic crosssectional measurement
Flow measuring
  Only in descending aorta 70% total flow

## Slide 97: Partial CO2 rebreathing technique
method applies Fick’s principle
by using expired carbon dioxide (CO2) concentration as an indicator
VCO2 - the difference between inspired and expired CO2 content
CvCO2 - estimated by using a partial rebreathing technique
CaCO2 is estimated from the PaCO2 and the end-tidal CO2

## Slide 98: Partial CO2 rebreathing technique
Because the diffusion rate of carbon dioxide is 22 times more rapid than that oxygen, it is assumed that no difference in venous CO2 (CvCO2) will occur

## Slide 102: Partial CO2 rebreathing technique
Advantages
  minimal invasiveness
  capability of continuous monitoring of CO
Limitations
  intubated patients without severe gas-exchange abnormality
  patients with PaCO2 above 30 mmHg
  patient who cannot tolerate a brief rebreathing period

## Slide 103: Non-invasive cardiac output monitoring

## Slide 104: Thoracic electrical bioimpedance (TEB)
delivery of a low-amplitude high-frequency electrical current across the thorax
sensing electrodes measuring impedance are placed on the upper and lower thorax
Hemodynamic parameters are measured based on changes in the thoracic electrical conductivity to changes of thoracic aortic blood flow during the cardiac cycle

## Slide 105: Thoracic electrical bioimpedance (TEB)

## Slide 106: Thoracic electrical bioimpedance (TEB)
TEB involves delivery of a low-amplitude high-frequency electrical current across the thorax
changes in the thoracic electrical conductivity to changes of thoracic aortic blood flow during the cardiac cycle
measuring the impedance change generated by the pulsatile flow and the time intervals between the changes, SV can be calculated

## Slide 107: Thoracic electrical bioimpedance (TEB)
Advantage
  TEB is a completely noninvasive CO monitoring method
Limitations
  arrhythmia
  fluid in the thoracic component
  noise from mechanical ventilation or surgical electrocautery

## Slide 108: Electrical bioreactance cardiography

## Slide 109: Electrical bioreactance cardiography(EB)
EB analysis is based on changes in frequency of electrical resistivity across the thorax
EB signal is less susceptible to interference from chest wall movement, lung edema, and pleural effusion. EB technology is commercially available as the NICOM system in the U.S.

## Slide 110: Electrical bioreactance cardiography
Advantages
  Noninvasive
  continuous monitoring with more variety in clinical applications (e.g., from small children to adults) and is very safe for clinical use

## Slide 111: Electrical bioreactance cardiography
Limitations
  signal interference was reported by electrocautery causing transiently impaired signals
  during episodes of low flow, NICOM signals may lose their accuracy

## Slide 112: The ccNexfin system
completely noninvasive continuous CO monitoring system
ccNexfin system provides various hemodynamic parameters including continuous BP, SV, CO, SVV, and SVR

## Slide 113: The ccNexfin system

## Slide 114: The ccNexfin system
Essence of the volume clamp method involves
  Clamping artery to a constant volume by dynamically providing equal pressure on either side of arterial wall
Volume - Measure by a photoplethysmograph built into cuff
Counter pressure - apply by an inflatable bladder inside the cuff and adjust 1000 times per second to keep arterial volume constant

## Slide 115: The ccNexfin system
Advantages
  noninvasive CO monitoring
Limitations
  requires the finger cuff to be inflated continuously
  maximum of 8 h per finger
  severe peripheral vasoconstriction
  edematous fingers
  regurgitant aortic valve
  aneurysm in the proximal aorta

## Slide 116
Device | Invasiveness | Equipment required | Limitations | Information
PAC- TD Continuous PAC_TD | ++++ | - Central venous access - Thermodilution capable PAC | PAC related complication | PAP,PCWP, SvO2
Pulse contour analysis PiCCO +TPTD | +++ | Thermistor-tipped arterial line in a central vessel CVC | -Requires intermittent recalibration and a central arterial catheter -Less accurate with significant aortic regurgitation, arrhythmia, or intra-aortic balloon pump | EVLW, GEDV, ITBV, SVV, PVV, ScvO2
Pulse contour analysis LiDCO +TPTD | +++ | Arterial line +/-CVC | -Requires intermittent recalibration -Intracardiac and extracardiac Shunts -Require good quality of arterial waveform -Calibration affected by muscle relaxant and lithium therapy | ITBV, SVV, PVV

## Slide 117
Device | Invasiveness | Equipment required | Limitations | Information
Volume View | +++ | Require A-line and CVP | --Require good quality of arterial waveform | EVLW, PVPI, GEDV, ITBV, GEF, SVV, PPV, ScvO2
floTrac/ Vigileo | ++ | A-line | - No external calibration -Arterial signal quality -Rapid changes in vascular motor tone -Not indicated for IABP - Inaccuracy in hemodynamic instability | CO,SVV, PPV, MAP
PRAM | ++ | A-line | - Arterial signal quality -Rapid changes in vascular motor tone - Inaccuracy in such an abnormal arteries | SVV, PPV, MAP

## Slide 118
Device | Invasiveness | Equipment required | Limitations | Information
TEE | + | TEE probe | Cost Operator dependent | EF, LVEDA, diameter of IVC/SVC
ED | + | ED probe | - Poorly tolerated unless tracheal tube present | CO,SV,PV
Partial CO2 rebreathing | + | NICO Rebreathing circuit ETT A-line | - Requires tracheal intubation and stable tidal volumes during measurements - accuracy limits in patient with abnormal V/Q mismatch -Valid only with CO2 >30 mmHg - patient who cannot tolerate a brief rebreathing period | CO

## Slide 119
Device | Invasiveness | Equipment required | Limitations | Information
volume clamp method | - | ClearSight system ccNexfin | -measurement is restricted to 8 h - strong vasoconstriction, Raynaud disease or very edematous fingers - erroneous may occur in patient with aortic valve insufficiency or proximal aneurysm | MAP, SVV, PPV
Bioimpedance | - | Biomed electrode | -movement artifacts, e.g., noise from mechanical ventilator, electrocautery -thoracic fluid overload -arrhythmia - need to be intubated Signal stability fails after 24 h | CO
Bioreactance | - | NICOM electrode | -limit in condition of low flow period -signal artifact, e.g., electrocautery | CO

## Slide 120: References
Minimally invasive or noninvasive cardiac output measurement:an update , Japanese Society of Anesthesiologists 2016, 30:461–480
Annals of Cardiac Anaesthesia , Vol. 11:1 ,Jan-June-2008
Minimally invasive cardiac output monitors,Continuing Education in Anaesthesia, Critical Care & Pain ,Volume 12 Number 1 2012
Newer methods of cardiac output monitoring, World Journal of Cardiology 2014 September 26; 6(9): 1022-1029
Invasive and non-invasive methods for cardiac output measurement, International Journal of Caring Sciences, 1(3):112–117